hOS

Regulatory Overview — Medical Device Classification & Approval Pathway

Australia (TGA) · International Expansion Strategy · March 2026

# 1. What hOS Is — Regulatory Classification

hOS is a local-first diagnostic interpretation platform that ingests data across multiple medical domains, runs deterministic evaluations, contextualises results via LLM, and synthesises cross-domain findings through a central Orchestrator.

From a regulatory standpoint, hOS is a Software as a Medical Device (SaMD) — software intended to be used for one or more medical purposes without being part of a hardware medical device.

| Attribute | Classification |
| --- | --- |
| Device type | Software as a Medical Device (SaMD) |
| Intended use | Diagnostic decision support for individuals |
| Risk tier (IMDRF) | Class IIb–III (serious condition, informing/driving diagnosis) |
| Australian framework | TGA — Therapeutic Goods Act 1989 |
| US framework | FDA 510(k) — Class II, product codes QIH / DQO |
| EU framework | EU MDR 2017/745 — Class IIa–IIb (Notified Body required) |

# 2. hOS as an Aggregator — What This Means Regulatorily

hOS is designed as an aggregator of already-cleared or validated ML tools. This is a meaningful regulatory advantage, but it does not fully exempt the platform from classification.

What the aggregator model protects:

- Individual node ML tools carry their own cleared status — hOS does not need to re-validate the underlying AI models

- The clinical validation burden is significantly reduced compared to building novel AI from scratch

- There is clear regulatory precedent (e.g. Aidoc aiOS) for aggregation platforms

What the aggregator model does NOT protect:

- The aggregation logic itself — combining outputs from multiple nodes into a unified view — is a novel medical device function requiring its own validation

- The Orchestrator's cross-diagnostic Evaluation layer is entirely hOS's own clinical logic and carries full regulatory responsibility

- LLM summarisation layers are not covered by any existing clearance and remain the most uncharted area for regulators globally

# 3. Architecture — Layer-by-Layer Regulatory Assessment

Each of the five pipeline layers carries a different regulatory footprint:

| Layer | Function | Regulatory Risk | Key Requirement |
| --- | --- | --- | --- |
| 1. Importing | Ingest raw data (PDFs, imaging, measurements) | None | HIPAA / Privacy Act compliance only |
| 2. Unifying | Normalise to canonical structured data | Very Low | Accuracy and auditability of mapping logic |
| 3. Evaluation | Deterministic analysis vs. clinical criteria — no LLM | HIGH — this is the regulated layer | Clinical criteria must be sourced from published guidelines or cleared tools; validation study required |
| 4. Insight (LLM) | Contextualise results — cannot override Evaluation | Medium | Must be technically isolated from clinical output; cannot contribute to diagnosis |
| 5. Pipeline | Route structured output to UI and Orchestrator | None | Data integrity and contract schema |

Critical architectural requirement:

The LLM constraint ('cannot override Evaluation output') must be technically enforced in the data model — not just architectural policy or prompt instruction. The clinical output pathway and the LLM explanation pathway must be structurally separate and independently auditable.

# 4. Node-by-Node Risk Assessment

| Node | Domain | Risk Level | Why |
| --- | --- | --- | --- |
| Anthropometry | Body composition | Low | Largely wellness territory; established metrics (BMI, body fat%). May be exempt from device classification depending on claims. |
| Hepatology | Blood work | Medium | Well-established reference ranges; clear predicates exist. Clinical stakes lower than imaging. Good starting node for approval. |
| Radiology | Medical imaging | High | Requires cleared ML tools underneath. Imaging is highest-scrutiny category. MRI specifically requires careful vendor selection. |
| Orchestrator | Cross-diagnostic synthesis | Highest | Novel function with no direct predicate. Cross-domain interaction rules require original clinical validation. Cannot leverage node clearances. |

# 5. International Regulatory Strategy

Australia's TGA has adopted the IMDRF SaMD framework — the same framework used by Health Canada, Singapore HSA, and aligned with FDA/EU MDR. This enables a 'clear once, expand many' strategy.

| Market | Regulator | Framework | Strategy |
| --- | --- | --- | --- |
| Australia | TGA | IMDRF-aligned | Start here — lower cost, accessible pre-submission consultation, strong international credibility |
| Canada | Health Canada | IMDRF-aligned | TGA clearance directly supports submission — reduced duplicated validation |
| Singapore | HSA | IMDRF-aligned | Strong APAC gateway; recognises IMDRF dossiers |
| USA | FDA 510(k) | Own framework | TGA dossier strengthens submission; product codes QIH or DQO. Target Year 4+ |
| EU / UK | CE Mark / UKCA | EU MDR 2017/745 | Most demanding — requires Notified Body; covers 27 countries. Target Year 4–5 |
| Japan | PMDA | Unique | Most complex in APAC; pursue only after other markets established |

# 6. Recommended Approval Roadmap

Clearance should be pursued incrementally by node, starting with lowest complexity. Each cleared node builds evidence for the next.

| Phase | Action | Est. Timeline | Est. Cost |
| --- | --- | --- | --- |
| Phase 0 Now | Engage TGA for free pre-submission consultation. Confirm Anthropometry classification (may be wellness-exempt). Establish data consent architecture. | 1–3 months | ~$0 |
| Phase 1 Year 1 | Implement ISO 13485 Quality Management System. Design clinical output contract schema between nodes. Enforce LLM/clinical layer separation technically. | 6–12 months | $50K–$150K |
| Phase 2 Year 1–2 | Clinical validation study for Hepatology (blood work) Evaluation layer. TGA submission for Hepatology node as first cleared feature. | 12–18 months | $150K–$300K |
| Phase 3 Year 2–3 | Integrate cleared radiology ML tools. Validate Radiology node aggregation layer. TGA submission for Radiology node. | 12–18 months | $200K–$400K |
| Phase 4 Year 3–4 | Validate Orchestrator cross-diagnostic logic with accumulated node evidence. TGA submission for full platform. Expand to Canada and Singapore. | 12–18 months | $150K–$250K |
| Phase 5 Year 4+ | FDA 510(k) and EU MDR submissions using full TGA dossier as foundation. | 12–24 months | $300K–$600K |

Total estimated investment to full TGA clearance (Phases 0–4): $550K–$1.1M over 3–4 years.

This is significantly lower than building novel AI diagnostic tools from scratch ($2M–$5M+), because the aggregator model offloads the heaviest validation burden to the underlying cleared tools.

# 7. Key Risks & Mitigations

| Risk | Severity | Mitigation |
| --- | --- | --- |
| LLM layer crosses into clinical decision territory | High | Technically enforce separation in data model; LLM output must never appear in clinical output schema |
| Orchestrator cross-diagnostic logic lacks clinical evidence | High | Engage clinical advisors early; ground interaction rules in published literature; treat as independent validation project |
| Underlying cleared tools are updated or decommissioned | Medium | Maintain vendor agreements with SLA on clearance status; architect nodes to be replaceable |
| Marketing language implies diagnosis before clearance | High | Legal review of all copy; use 'for informational purposes' framing until cleared |
| QMS not in place before building critical features | Medium | Establish ISO 13485 QMS in Phase 1 — retrofitting is significantly more expensive than building in |
| EU MDR Notified Body bottleneck | Low (long term) | Engage Notified Body early; EU MDR timelines are 18–36 months minimum |

# 8. Immediate Next Steps

- Book a TGA pre-submission consultation (free) — confirm classification tier for each node and Orchestrator

- Engage a regulatory affairs consultant with SaMD experience in Australia

- Audit current architecture against ISO 13485 requirements — identify gaps before further development

- Define the clinical output contract schema between nodes and Orchestrator formally

- Review all marketing and UI copy for language that could imply clinical diagnosis

- Identify cleared radiology ML tool vendor for Radiology node integration

This document is a strategic overview for internal planning purposes only and does not constitute legal or regulatory advice. Engage a qualified regulatory affairs consultant before making submissions to TGA or any other authority.